INTERLABORATORY COMPARISONS OF CALIBRATION RESULTS OF THREE-PHASE ELECTRICAL ENERGY METERS
Oleh Velychko, Dr.Sc., Prof., Yulia Kulish, Serhii Kursin, PhD, 
State Enterprise “Ukrmetrteststandard”, Ukraine,
Boris Gendelman,
SATEC Calibration Lab., Israel

e-mail: velychko@hotmail.com
Abstract

Reliable measurement of electric energy is of great importance both in industry and for household consumers of both single-phase and three-phase networks. Interlaboratory comparisons (ILCs) are one of the experimental forms of confirming the scope of accreditation of calibration laboratories (CLs). The article is devoted to the analysis of the results of the ILC of a three-phase electric energy meter, which was conducted in 2025. The ILC was conducted with the participation of accredited CLs from Ukraine and Israel. A feature of this ILC was the calibration of the meter both for three phases in general and for each phase separately. Calibration was carried out with a nominal AC voltage of 230 V, frequency 50 Hz, currents from 2.5 A to 100 A at KP 1.0 and 0.5 for positive and negative flow. Processing of the primary ILC data received from the laboratories was carried out using special software “InterLab 1.0”, developed in UMTS. The deviation of the results obtained by each laboratory was determined and the consistency of the results obtained was assessed taking into account the measurement uncertainty using the criterion of performance statistics. In general, the laboratories obtained satisfactory results with good consistency. The laboratories fully meet the established requirements for the En number and have confirmed their technical competence during the calibration of meters in accordance with the requirements of the international standard ISO/IEC 17025. The results of the ILC allow for the mutual recognition of the results of the calibration of electricity meters by the national accreditation bodies that accredited these CLs.
Keywords

Calibration, energy meter, three-phase voltage, interlaboratory comparison, uncertainty.
1. Introduction

Reliable measurement of electrical energy is important both in industry and for household consumers of both single-phase and three-phase networks. Such measurements affect the accurate calculations of the cost of consumed electrical energy, allow you to manage its consumption and diagnose the electrical network, etc. Monitoring electrical energy consumption helps to detect overloads or malfunctions in the network.
National agencies for laboratory accreditation establish specific requirements for the validation of the scope of accreditation of calibration laboratories (CL) [1] for each type of measurement and range of measured value. Interlaboratory comparisons (ILCs) [2] are one of the experimental forms of such validation of the scope of accreditation of a CL. They are used by accredited calibration laboratories (CLs) to assess and confirm the accuracy, reliability and reproducibility of measurements carried out by these laboratories.

The main purpose of ILCs is to assess the accuracy of measurement standards and measuring instruments used in the laboratory, to confirm the mutual acceptability of results obtained by laboratories to increase confidence in their results, etc. Successful results for a particular laboratory in ILCs are confirmation of its competence in carrying out certain types of measurements by a specific specialist using specific standards and measuring instruments.
2. Drawbacks

A large number of publications are devoted to the issues of organizing ILC and methods of processing the obtained data in specific types of measurements or tests. The assessment of the laboratory results obtained in ILC for specific types of measurements is presented in a number of works, in particular on electrical measurements (electrical energy, reactive power, electrical resistance) [4–8]. The assessment of the measuring capabilities of CL for obtaining high-accuracy and precision data is given in the works [9–11]. At the same time, the actual task of conducting ILC for three-phase electricity meters with the appropriate assessment of its results remains.
3. Goal

The purpose of the study was to evaluate the primary data of laboratories obtained during the ILC for the calibration of three-phase electricity meters. During the study, it was necessary to examine the calibration object in the corresponding ILC, determine the assigned value and its expanded uncertainty, and calculate the degree of equivalence for each of the laboratories participating in the ILC and their expanded uncertainties. The results obtained by the laboratories in the ILC had to be evaluated according to the criteria of performance statistics with the formulation of the corresponding conclusions.
4. Overview of the interlaboratory comparison
The ILC for the calibration of a three-phase electricity meter (UMTS ILC-E5:2025) was conducted by two CLs (SE “Ukrmetrteststandart” – UMTS, Ukraine and SATEC Calibration Laboratory, Israel) from July to September 2025. The UMTS laboratory is accredited by the National Accreditation Agency of Ukraine (NAAU), and the SATEC laboratory is accredited by the Israel Laboratory Accreditation Authority (ISRAC). Both NAAU and ISRAC are signatories to the International Accreditation Cooperation (ILAC) agreement on mutual recognition of calibration results.
The ILC program was developed with requirements of international standards ISO/IEC 17025 [1], ISO/IEC 17043 [2] and ISO 13528 [3].
The calibration item (CI) for ILC is the three-phase energy meter SATEC EM133-XM-HACS(SE) 3×50 A (100 A), serial number of which is 75000208 (Fig. 1). SATEC EM133 meters are a whole family of energy meters, which includes the electricity meter selected for comparison. It has a nominal measured current of 3×50 A (maximum current 100 A) and is intended for applications with transformer power supply and connection to high-voltage power lines.
[image: image1.jpg]SSSATEC

EM133-XM-HACS(SE
Class 0.58 Op. temp:-20C to 60C
3 x 230/400 V_50Hz/60Hz
Rated Current \npu! In (Imax)

T €
BLOBOY REY.A1 0" >

11 §Z LetgELd |3 neTToo
fb 10 11 2+ , X X





Fig. 1. Appearance of EM133-XM-HACS(SE) 3×50 A (100 A)
Main characteristics of the CI for ILC:

– measured voltage/supply voltage



from 120/207 V AC to 230/400 V AC (L-N/L-L);

– maximum current




100 A;

– frequency measurement range



50/60 Hz;

– rated currents of the remote current transformer

40 mA

– voltage load for 277 V




< 1.5 VA;

– operating temperature range



from -20°C to 60°C;

– dimensions





125 × 90 × 75 mm.

Main measurements should be performed with the input signals and environmental conditions:

– AC voltage





230 V ± 0.05 %;

– current






from 2.5 to 100 A ± 0.05 %;

– power factor
 (PF)




1.0, ±0.5 Lag;

– frequency





(50 ± 0.01) Hz;

– ambient temperature




(23 ± 2) °С;

– relative humidity




(40 ± 10) %;

– supply voltage





(230 ± 3) V;

– frequency of the supply voltage



(50 ± 0.05) Hz.
Traceability to the International System of Units (SI) was ensured by calibrating the reference standards used by the participating laboratories in accredited calibration laboratories (Fig. 2). Electrical energy measurements in the UMTS laboratory have metrological traceability to the Physikalisch-Technische Bundesanstalt of Germany (PTB), and SATEC calibration laboratories are traceable to the Czech Metrology Institute (CMI). The UMTS laboratory was a pilot laboratory for the verification of electrical power and energy standards of two regional metrology organizations: the key power verification COOMET. EM-K5 [13] and the additional energy verification GULFMET.EM-S5 [14], which also confirm the traceability of the corresponding measurements.
[image: image2.jpg]



Fig. 2. Appearance of the National Standard of Ukraine for Units of Electric Power and Power Factor
NDETU EM-08-2023
5. The obtained results of the interlaboratory comparison
The full final report of ILC contains all the primary data of the measurements of total electrical energy for all phases, as well as for each phase separately, and the corresponding uncertainty estimates of the ILC participating laboratories. The report includes a description of the measurement method, traceability to SI, as well as the results and associated uncertainties. The calibration errors xi and their expanded uncertainties U(xi) of the participating laboratories are given in Table 1 for an AC voltage of 230 V at a frequency of 50 Hz.
Table 1. Results of measurement of total electrical energy across all phases by ILC participating laboratories
	Current,
A
	PF
	Flow Positive/ Negative
	Calibration error xi,
%
	Uncertainty U(xi),
%

	
	
	
	UMTS
	SATEC
	UMTS
	SATEC

	L1

	2.5
	1
	Positive
	0.016
	0.012
	0.115
	0.147

	5
	
	
	0.045
	0.023
	0.106
	0.131

	50
	
	
	0.016
	-0.008
	0.176
	0.175

	100
	
	
	0.031
	-0.026
	0.097
	0.160

	5
	0.5 Lag
	
	0.150
	0.230
	0.091
	0.334

	50
	
	
	0.017
	-0.008
	0.161
	0.236

	100
	
	
	-0.023
	-0.001
	0.060
	0.245

	50
	1
	Negative
	-0.011
	-0.013
	0.129
	0.168

	100
	
	
	-0.037
	-0.029
	0.048
	0.163

	50
	0.5 Lag
	
	0.079
	0.034
	0.179
	0.257

	100
	
	
	-0.069
	-0.073
	0.160
	0.269

	L2

	2.5
	1
	Positive
	0.008
	0.019
	0.116
	0.147

	5
	
	
	0.047
	0.044
	0.105
	0.131

	50
	
	
	0.016
	-0.002
	0.173
	0.175

	100
	
	
	0.033
	-0.014
	0.100
	0.160

	5
	0.5 Lag
	
	0.089
	0.199
	0.089
	0.334

	50
	
	
	-0.040
	0.044
	0.164
	0.237

	100
	
	
	-0.081
	-0.008
	0.060
	0.245

	50
	1
	Negative
	-0.010
	-0.003
	0.129
	0.168

	100
	
	
	-0.034
	-0.018
	0.047
	0.163

	50
	0.5 Lag
	
	0.027
	0.011
	0.179
	0.257

	100
	
	
	-0.127
	-0.022
	0.158
	0.270

	L3

	2.5
	1
	Positive
	-0.015
	0.000
	0.117
	0.147

	5
	
	
	0.047
	0.042
	0.104
	0.131

	50
	
	
	0.010
	-0.002
	0.175
	0.175

	100
	
	
	0.023
	-0.021
	0.101
	0.160

	5
	0.5 Lag
	
	0.112
	0.099
	0.089
	0.334

	50
	
	
	-0.025
	-0.049
	0.161
	0.236

	100
	
	
	-0.061
	-0.056
	0.059
	0.245

	50
	1
	Negative
	-0.017
	-0.004
	0.127
	0.168

	100
	
	
	-0.043
	-0.017
	0.048
	0.163

	50
	0.5 Lag
	
	0.039
	-0.067
	0.178
	0.257

	100
	
	
	-0.121
	-0.087
	0.159
	0.269

	Three phases

	2.5
	1
	Positive
	0.027
	0.011
	0.034
	0.116

	5
	
	
	0.059
	0.037
	0.034
	0.095

	50
	
	
	0.023
	-0.003
	0.034
	0.123

	100
	
	
	-0.007
	-0.021
	0.038
	0.118

	5
	0.5 Lag
	
	0.206
	0.175
	0.034
	0.259

	50
	
	
	0.012
	-0.005
	0.033
	0.194

	100
	
	
	-0.060
	-0.021
	0.035
	0.196

	50
	1
	Negative
	-0.060
	-0.021
	0.035
	0.196

	100
	
	
	0.013
	-0.004
	0.036
	0.119

	50
	0.5 Lag
	
	-0.016
	-0.022
	0.045
	0.120

	100
	
	
	0.031
	0.072
	0.033
	0.202


Laboratories calculated their measurement uncertainty according to the Guide to the Expression of Uncertainty in Measurement (GUM) [15]. All contributions to the measurement uncertainty were reported in the Final Report of ILC and identified as Type A or Type B uncertainties. The total standard uncertainty was calculated from all individual standard uncertainties. The expanded uncertainty was calculated with a coverage factor of 2 at a confidence level of 95%.
6. Discussion of the results of the interlaboratory comparison
The ILC assigned values (AV) XAV are calculated as the mean of participant data:


[image: image3.wmf](

)

(

)

22

11

1

,

NN

i

AV

ii

ii

x

X

uxux

==

=

åå






(1)

with expanded uncertainties


[image: image4.wmf](

)

(

)

22

UMTSSATEC

11

()21,

AV

UX

uxux

æö

=+

ç÷

èø





(2)
where xUMTS and xSATEC are measurement error for UMTS and SATEC accordingly, u(xUMTS) and u(xSATEC) are combined standard uncertainty for UMTS and SATEC accordingly.

AV with expanded uncertainties for three phases is given for AC voltage 230 V and at frequency of 50 Hz in Table 2.

Table 2. AV and expanded uncertainties.
	Current,
A
	PF
	Flow Positive/ Negative
	AV XAV,

%
	Uncertainty U(XAV),

%

	2.5
	1
	Positive
	0.026
	0.0326

	5
	
	
	0.057
	0.0320

	50
	
	
	0.021
	0.0328

	100
	
	
	-0.008
	0.0362

	5
	0.5 Lag
	
	0.205
	0.0337

	50
	
	
	0.012
	0.0325

	100
	
	
	-0.059
	0.0345

	50
	1
	Negative
	0.012
	0.0345

	100
	
	
	-0.017
	0.0421

	50
	0.5 Lag
	
	0.032
	0.0326

	100
	
	
	-0.069
	0.0345


The degrees of equivalence (DoE) of i-th laboratory with expanded uncertainties with respect to the AV is estimated as


[image: image5.wmf]  

,

labilabiAV

DxX

=-







(3)

[image: image6.wmf](

)

(

)

22

  

().

labilabiAV

UDUxUX

=+






(4)
DoE with expanded uncertainties for laboratories for three phases for AC voltage 230 V at frequencies of 50 Hz are given in Table 3. In the first column of this table, each AC value is assigned a number that can be used to display the DoE of laboratories in a graphical form (Fig. 3). To analyze the ILC results, special software “InterLab 1.0” [16] developed by UMTS was used, which makes it possible to automate the calculations and analysis of the ILC results evaluation criteria. It allows evaluating the ILC results by the En number, as well as other similar criteria. The obtained results of the deviation and uncertainty estimates are displayed using graphs, and the calculated values of the indicators are displayed using histograms.
Table 3. DoE with expanded uncertainties for laboratories.
	Current (No),
A
	PF
	Flow Positive/ Negative
	DoE 
[image: image7.wmf] 

labi

D

,
%
	Uncertainty of DoE 
[image: image8.wmf] 

()

labi

UD

,
%

	
	
	
	UMTS
	SATEC
	UMTS
	SATEC

	2.5 (1)
	1
	Positive
	0.001
	-0.015
	0.0471
	0.1205

	5 (2)
	
	
	0.002
	-0.020
	0.0467
	0.1002

	50 (3)
	
	
	0.002
	-0.024
	0.0472
	0.1273

	100 (4)
	
	
	0.001
	-0.013
	0.0525
	0.1234

	5 (5)
	0.5 Lag
	
	0.001
	-0.030
	0.0479
	0.2612

	50 (6)
	
	
	0.000
	-0.017
	0.0463
	0.1967

	100 (7)
	
	
	-0.001
	0.038
	0.0491
	0.1990

	50 (8)
	1
	Negative
	0.001
	-0.016
	0.0498
	0.1239

	100 (9)
	
	
	0.001
	-0.005
	0.0616
	0.1272

	50 (10)
	0.5 Lag
	
	-0.001
	0.040
	0.0464
	0.2046

	100 (11)
	
	
	0.000
	0.009
	0.0491
	0.2089


[image: image9.png]Comparison results

1 T1031vS

TISLAN

i 01231¥S

oTSLWN

[REIS

6sLN

EREINS

8sLN

i 1231vs

Lsinn

i 9231vs

- 9S1AN

-t soauvs

ssLwn

[eEIS

pSLNN

[REITS

€SN

2031V

zsLnn

1031V

TSLAN

S
T

faureyzoun/ious Jo anjen

9
S

02

01
-02
-03

Laboratories




Fig. 3. DoE with expanded uncertainties for laboratories.

The data consistency En number is calculated as:


[image: image10.wmf](

)

    

21.0.

nlabilabilabi

EDUD

=£






(5)
En number for laboratories for three phases and AC voltage 230 V at frequencies of 50 Hz are given in Table 4. In the first column of this table, each AC value is assigned a number that can be used to display the En number of laboratories in a graphical form (Fig. 4). Laboratories meet the established requirements for En number (ǀEnǀ ≤ 1.0): UMTS – from 0.01 to 0.47; SATEC – 0.00 to 0.50 for all calibration points.

Table 4. En number for laboratories for three phases.
	Current (No),
A
	PF
	Flow Positive/ Negative
	En for UMTS
	En for SATEC

	2.5 (1)
	1
	Positive
	0.03
	0.12

	5 (2)
	
	
	0.05
	0.19

	50 (3)
	
	
	0.04
	0.19

	100 (4)
	
	
	0.03
	0.10

	5 (5)
	0.5 Lag
	
	0.01
	0.12

	50 (6)
	
	
	0.01
	0.08

	100 (7)
	
	
	0.02
	0.19

	50 (8)
	1
	Negative
	0.03
	0.13

	100 (9)
	
	
	0.01
	0.04

	50 (10)
	0.5 Lag
	
	0.02
	0.20

	100 (11)
	
	
	0.01
	0.04


[image: image11.jpg]Quantitative indicators "EnXav2"

98

0.183

"

0.20
015
0.10

0.05
0.00 4

0
e
S
T
103e21pUl 3 JO BNjeA BYL

e
=

-0.15

-0.20

T13Lvs

TISLWN

0T23LYS

0TSLAN

RIS

6sLN

8031vS

8sLN

JeEIS

Lsinn

EREIS

9sLIN

RIS

ssLwn

$23LVS

pSLNN

[REITS

€sLnn

2031V

zsinn

QRIS

TSLWN

Laboratories




Fig. 4. En number for laboratories for three phases.

7. Conclusions

The ILC on the calibration of three-phase electricity meters was carried out with the participation of accredited CLs from Ukraine and Israel. A feature of this ILC was the calibration of the meter both for three phases in general and for each phase separately. Calibration was carried out with a nominal value of alternating voltage of 230 V with a frequency of 50 Hz at currents from 2.5 A to 100 A at PF 1.0 and 0.5 for positive and negative flow. Processing of the primary ILC data received from the laboratories was carried out using special software “InterLab 1.0”, developed at UMTS. In general, the laboratories received satisfactory results with good consistency. The laboratories fully comply with the established requirements for the En number (ǀEnǀ ≤ 1.0) and have confirmed their technical competence in calibrating meters in accordance with the requirements of the ISO/IEC 17025 standard. The results of the ILC allow for mutual recognition of the calibration results of electricity meters by the national accreditation bodies that accredited these CLs.
Gratitude 

The authors express their gratitude to the National Research Foundation of Ukraine, project number 0123U103529 (2022.01/0009) “Assessing and forecast- ing threats to the reconstruction and sustainable opera- tion of objects of critical infrastructure” from the contest “Science for the reconstruction of Ukraine in the war and post-war periods”.
 Conflict of Interest 

The authors state that there are no financial or other potential conflicts regarding this work.
References
Example for Materials of Conference

[1] M. Mykyychuk, Yu. Yatsuk, O. Ivakhiv, R. Matviiv, “Voltage and Resistance Calibrators for Verification of Industrial Instrument Applications”, in Proc. Metrol. Com. of Katowice branch of Pol. Acad. Sciences. Series: Conf. no. 21, 12th Conference “Problems and Progress in Metrology”, Szczyrk, Poland, 2016, pp.114-117. 
Example for Article

[2] V. Yatsuk, M. Mykyychuk, “Remote Errors Correction of Multi-Channel Cyber-Physical Measuring Systems”, Adv, Cyber-Physical Systems, Lviv Polytech. Publ. House, no.1, p.16-21, 2016. 
Example for Book
 [3] G. Mejer, Smart Sensor Systems. John Wiley, 2008. 
Example for Online information
[4] Fluke Multifunction Calibration Tools, Fluke Inc., 2015. [Online]. Available: http://en-us.fluke.com/products/multifunction-calibrators/.
Example for Patent
[5] B. Stadnyk, S.Yatshyshyn, “A method for obtaining a temperature quantum based on fundamental constants of the matter and establishing a temperature standard and a device for its implementation”, G01N 27/00, Pat.115601 UA, bul.24, 27.11. 2017.
Example for Conference Series
[6] P. Ripka, “Security applications of magnetic sensors”, Journ. Phys., Conference Series, no.6, 2013, 450(1):2001-, doi:10.1088/1742-6596/450/1/012001.
Example for Standards
[7] DSTU-N RMH 43:2006. Metrolohiia. Zastosuvannia “Rukovodstvj po vyrazhenyiu neopredelennosty yzmerenyi” (RMH 43:2001, IDT).

[8] ISO 9000:2015 Quality management systems -- Fundamentals and vocabulary , 2015.  
!!!Використана література повинна бути не старшою 2020 року.

Кількість самопосилань 10%.

Використаної літератури повинно бути не менше 10-15 пунтів.

ISSN: 2617-846X (online)

 Velychko Oleh
Number ORCID: https://orcid.org/0000-0002-6564-4144
Kulish Yuliia
Number ORCID: https://orcid.org/0000-0002-5041-8398
Kursin Serhii
Number ORCID: https://orcid.org/0000-0003-2147-6211
Geidelman Boris
Number ORCID: -
Date of receipt of the article -08.10.25.

Date of acceptance after review – 20.11.25.

Date of publication of the article- 30.12.25.

Data about authors

Яцишин Святослав Петрович 
Доктор технічних наук, професор 

Кафедра «Інформаційно-вимірювальні технології»

Національний університет «Львівська політехніка»

вул. С. Бандери, 12, м. Львів, Україна, 79013

Е-mail: yatsuk.vasyl@gmail.com

Контактний тел.: +38 067 30 34 830

Кількість статей у загальнодержавних базах даних – 90

Кількість статей у міжнародних базах даних – 19

Номер ORCID: https://orcid.org/0000-0002-4213-4862 

Svyatoslav Yatsyshyn 
Doctor of Technical Sciences, Professor 

Department of Information Measurement Technologies

Lviv Polytechnic National University
Bandera str., 12, Lviv, Ukraine, 79013

Е-mail: yatsuk.vasyl@gmail.com

Contact tel.: +38 067 30 34 830

The number of articles in national databases – 90

The number of articles in international databases – 19

Number ORCID: https://orcid.org/0000-0002-4213-4862
Далі представити за аналогією інформацію про кожного співавтора
_1821417384.unknown

_1821425761.unknown

_1825147398.unknown

_1821417390.unknown

_1821425752.unknown

_1821417385.unknown

_1821417379.unknown

